## Slide 1: Reversible Stemless Shoulder Implant
Team 102
03/28/24

## Slide 2: Team Introductions
Angelina Lanh
Biomedical Engineer
Santiago Lazarte
Biomedical Engineer
John Sorensen
Biomedical Engineer
Taylor Vanderlinden
Biomedical Engineer
Kiersten Cady
Mechanical Engineer
Exactech Team
William Wartman
Mechanical Engineer
Kiersten Cady
Notes: Hello everyone, this is team 1x2 working with Exactech

## Slide 3: Sponsor and Advisors
Engineering Mentor
Tom Vanasse
Liaison Engineer at Exactech
Academic Advisor
Shayne McConomy, Ph.D.
Teaching Faculty II
Academic Advisor
Stephan Arce, Ph.D.
Teaching Faculty III
Kiersten Cady

## Slide 4: Objective
The objective of this project is to improve upon Exactech's
current reversible stemless shoulder implant, the Equinox
Kiersten Cady
Notes: The objective of this project is to improve upon Exactech's current reversible stemless shoulder implant, the Equinox. So, what is Equinox and how did I improve it? Before I can answer this question, some background information is needed.

## Slide 5: Background
Shoulder joint failure is common
Least stable joint in the body
Aging population
Current Models
  Stemmed: substantial bone loss
  Stemless: less stability but less bone loss
  Natural: greater range of motion
  Reverse: preferred for arthritis or rotator cuff damage
Kiersten Cady

## Slide 6: Stemmed vs Stemless
Stemmed shoulder implant
Stemless shoulder implant
Shoulder
Arm
Shoulder
Arm
Kiersten Cady

## Slide 7: Natural vs Reverse
Lazarte
Shoulder
Arm
Kiersten Cady
Notes: The reverse part of the reverse stemless arthroplasty is shown here on the left, versus the anatomical arthroplasty on the right. The anatomical model mimics how the actual shoulder joint is shaped, with the “ball” of the “Ball and Socket” joint on the humeral head as you can see with the dome-like implant head, and the socket on the scapula side at the glenoid.
In the reverse, the ball and socket positions are switched with the ball on the glenoid and the socket on the humeral head.
The reverse arthroplasty elongates the shoulder, allowing for a slightly larger range of motion. It also has a fixed fulcrum due to the offset and only requires the deltoid to be intact (the elongated shoulder stretches the deltoid, adding extra force that may have been needed from the rotator cuff).
In comparison, the anatomical model requires both he deltoid and the rotator cuff to be intact. Similar to a normal shoulder, the rotator cuff is responsible for keeping the joint tightly together, and the deltoid is the muscle in charge of movement.

## Slide 8: Requirements
Customer Needs
  Increase the strength of the interface between bone and implant
  Designs can be rigorously tested so that they exceed the current design in longevity and resistance to forces and moments
Stability
  Resist Forces - 10% improvement
Ergonomics
  Ease Implantation
Adaptability
  Variable Implant Sizes
Kiersten Cady

## Slide 9: Concepts
Concept Generation and Selection
  3 fin design (7.334 in^2) – Equinox Equivalent
  5 fin design (8.369 in^2)
  7 fin design (9.530 in^2)
Increase in surface area interaction makes for a stronger interface
Odd number of fins results in only one fin aligning with any force vector thereby increasing resistance while also limiting amount of additional bone loss
Kiersten Cady

## Slide 10: 3 Fin Design
William Wartman

## Slide 11: 5 Fin Design
William Wartman

## Slide 12: 7 Fin Design
William Wartman

## Slide 13: Even vs Odd #Fins
William Wartman
Force
Notes: 2D Sketch Projection of the 3 fin model

## Slide 14: Even vs Odd #Fins
William Wartman
Force
Notes: 2D Sketch Projection of the 3 fin model

## Slide 15: In Line vs Not in Line
Force
Force
William Wartman

## Slide 16: In Line vs Not in Line
3 Fin not inline 30 deg
5 Fin not inline 18 deg
7 Fin not inline 12.9 deg
William Wartman

## Slide 17
William Wartman
Bone Block
Implant
Force Gauge

## Slide 18: Test Stand Operation
William Wartman

## Slide 19: Update
Testing Apparatus
Implant Testing
  15 Tests Conducted
Data Processing
Matlab Code
Unusable Data
  Testing Rig Malfunctions
  Data Storage Issues
  Bone Block Cracking
Kiersten Cady

## Slide 20: 3 Fin 10 PCF
Kiersten Cady

## Slide 21: 3 Fin 10 PCF
Kiersten Cady
In Line
  Failure point determined at a maximum force of 9.2 lbf and 4.1 inches
  Withstood a larger force but failed at a smaller displacement point
  Early signs of failure or experimental error
Max Force

## Slide 23: 3 Fin 10 PCF
Kiersten Cady
Not In Line
  Failure point determined at a maximum force of 8.5 lbf and 4.4 inches
  Withstood a smaller force but failed at a larger displacement point
  Further testing to verify results
Max Force

## Slide 24: 5 fin 10 PCF
Kiersten Cady

## Slide 25: 5 fin 10 PCF
Kiersten Cady
Max Force
In Line
  Failure point determined at a maximum force of 11.1 lbf and 3.9 inches
  Withstood a larger force and larger displacement

## Slide 26: 5 fin 10 PCF
Kiersten Cady
Max Force
Not In Line
  Failure point determined at a maximum force of 9.6 lbf and 3.6 inches
  Withstood a smaller force and smaller displacement
  Further testing to verify results and examine errors

## Slide 27: 3-5 Fins 10 pcf
Inline
  3 fin 9.2 lbf
  5 fin 11.1 lbf
  20.6% Improvement
Not inline
  3 fin 8.5 lbf
  5 fin 9.6 lbf
  12.9% Improvement
William Wartman

## Slide 28: Incomplete Work
Additional 3 and 5 fin tests
7 Fin Testing
Fatigue Testing
William Wartman

## Slide 29: Lessons Learned
Biomedical Engineering
Bone Structures
Implants
Joints (mech sys 1)
Integrated Circuits
ASTM F-1839-08 Standard for Ortho foam blocks
William Wartman
Notes: “Standard Specification for Rigid Polyurethane Foam for Use as a Standard Material for Testing Orthopaedic Devices and Instruments”

## Slide 30: References
Equinoxe reverse system. www.exac.com. (2022, July 6).
https://www.exac.com/extremities/equinoxe-reverse-system/
Total shoulder replacement surgery. Shoulder Innovations. (2022, October 31). https://shoulderinnovations.com/total-shoulder-replacement-surgery/
McClintock, Dr. K. (2023, August 2). Pros and cons of reverse shoulder replacement surgery. Kyle McClintock, DO | Orthopedic Surgeon Roseville, CA. https://drmcclintock.com/pros-and-cons-of-reverse-shoulder-replacement-surgery/
Shoulder bones. Joint Preservation Center. (2013, April
5). http://josephbermanmd.com/anatomy-of-the-shoulder/shoulder-bones/
William Wartman

## Slide 31: Test Stand Assembly
William Wartman

## Slide 32
William Wartman

## Slide 33: Bending Stress
William Wartman
Notes: We don’t care about bending stress in the beam (tube) we only care about bending stress at the interface

## Slide 34: Furthering Failure Criterion
Similar to 0.2% offset in stress-strain graph
1st order linear fit with 0.2 displacement offset from the origin
Small dips likely due to accidentally contact with the shutoff switch during testing
William Wartman
Notes: Graph is of 10 PCF bone block using the 3 FIN model Block 8

## Slide 35: PuTTY Serial Output
Displacement (in)
Time (milliseconds)
Force (lbf)
William Wartman

## Slide 36: Testing Implantation Methods
Pressing Force
2
William Wartman
Bone Block

## Slide 37: Equinox
William Wartman

## Slide 38: Backup Slides

## Slide 39: Functional Decomposition
Reversible Stemless Shoulder Implant
Stability
Ergonomics
Adaptable
Kiersten Cady

## Slide 40: Functional Decomposition
Stability
Hold shoulder
Resist Shear Forces
Resist Rocking Moments
Resist Torque
Ergonomics
Easier to Implant
Adaptable
Variable Sized Implants
Kiersten Cady

## Slide 41: Targets and Metrics
Preventing premature implant failure
Resist shear, torque, or rocking moments
Improvements the ease installation
Kiersten Cady

## Slide 42: Targets and Metrics
Targets
Resists rocking moments at least 10% better than Equinox
Resists torque at least 10% better than Equinox
Resists shear forces at least 10% better than Equinox
Implantation is at least 10% better than Equinox
Kiersten Cady

## Slide 43: Targets and Metrics
Resist Rocking Moments
Moment (Newton-meter)
Resist Shear
Stress (Newton per square meter)
Resist Torque
Easier to Implant
Physical Exertion
Time
Steps
Torque (Newton-meter)
Kiersten Cady

## Slide 44: Targets and Metrics
Lifespan
Cost of Manufacturing
Recyclability
90% recyclable
Material composition and overall part density
10-20 years with minimal surgical revisions
Kiersten Cady

## Slide 45: Concept Generation
Anti-Problem
Battle of Perspectives
Biomimicry
Material Composition
Fixation Methods
Cylinder Geometry
Fin Geometry
Kiersten Cady

## Slide 46: Concept Generation
Medium Fidelity Concepts
Large diameter hollow cylinder with internal fins
Large diameter hollow cylinder with internal cylinder
Small diameter hollow cylinder with four external fins
Small diameter hollow cylinder with five external fins
Small diameter hollow cylinder with internal fins and external fins
Kiersten Cady

## Slide 47: Customer Needs
Question | Customer Statement | Interpreted Need
What are the current design’s causes of failure? | Rocking and shear stress. | Increase strength of the interface between bone and implant.
What do you want us to do? | Improve upon the current Exactech’s design. | Development of designs that can be rigorously tested such that they exceed the current design in longevity and resistance to certain forces and moments.
Would fasteners work to secure the implant to the bone? | No, due to the density of bone fasteners would not work. | Other methods of attaching are needed to prevent bone fragmentation.
How should be align implants for testing? | Last year's team made 3D printed punches, you will need something similar. | Design and implement punches to properly align implants with the substitutive bone surface.

## Slide 48: Customer Needs continued...
Question | Customer Statement | Interpreted Need
Does a current method of testing, these designs exist? | Last year's team focused on testing their designs using their own methods, you will need to do something similar. | We will need to develop a solid method of testing designs for max shear and bending moment.
What are some ways to imitate bone for implant testing? | Bone block or bone foam. | Bone block or bone foam.
Which factors are important besides the implant bone interface? | Mobility and preventing scapular notching are both important factors. | Our design must consider these two phenomena when designing.
What different configurations are currently available for these implants? | Both inlay and onlay, I would like to see you test both. | Investigate both and compare the results.

## Slide 49: Customer Needs continued...
Question | Customer Statement | Interpreted Need
Can cement or epoxy be used to secure the implant to the bone? | No due to revision surgeries. The cement and implant can't always be removed. | No due to revision surgeries. The cement and implant can't always be removed.

## Slide 50: Functional Decomposition
Function​ | Stability​ | Ergonomics​ | Adaptable​
Resist Rocking moments​ | X​ | ​ | X​
Resist Shear forces​ | X​ | ​ | X​
Resist Torque​ | X​ | ​ | X​
Easier to Implant​ | ​ | X​ | ​
Variable Sized Implants​ | ​ | X​ | X​

## Slide 51: Cup Design
William Wartman

## Slide 52: Cup Internal Cylinder Design
William Wartman

## Slide 53: Cup Internal 3 Fin Design
William Wartman

## Slide 54: HOQ
William Wartman

## Slide 55: Top Five Most Important Engineering Characteristics
Rank Order | 1 | 2 | 3 | 4 | 5
Engineering Characteristic | Rocking Moment | Fatigue | Shear | Easy to Manufacture | Overall Volume
William Wartman

## Slide 56: Pugh Chart
William Wartman

## Slide 57: Pugh Chart Continued
William Wartman

## Slide 58: Final Rating Matrix
Cup
6 Fin
Cup Internal Cylinder
William Wartman

## Slide 59: Inlay vs Onlay

## Slide 60: College of Engineering Color Palette
Fang
Orange
RGB: 238, 118, 36
Hex: #EE7624
CMYK:2, 66, 99, 0
1 2 3 4
1 2 3 4
White
Pantone: PMS 000C
RGB: 255, 255, 255
Hex: #FFFFFF
CMYK: 0, 0, 0, 0
1 2 3 4
Black
Pantone: Black C
RGB: 0, 0, 0
Hex: #000000
CMYK:0, 0, 0, 100
Garnet
Pantone: PMS 195 C
RGB: 120, 47, 64
Hex: #782F40
CMYK:19, 90, 50, 55
1 2 3 4
Notes: Avoid 75% Black in B135

## Slide 61: Accent Color Palette
1 2 3 4
Tardis Blue
RGB: 0, 59, 111
Hex: #003B6F
CMYK:
1 2 3 4
Turquoise
RGB: 64, 224, 208
Hex: #40E0D0
CMYK:
1 2 3 4
American Orange
RGB: 255, 139, 0
Hex: #FF8B00
CMYK:
Rubine Red
RGB: 206, 0, 88
Hex: #CE0058
CMYK: 0, 100, 43, 12
1 2 3 4
1 2 3 4
Asagi-iro
RGB: 72, 146, 155
Hex: #48929b
CMYK:
1 2 3 4
Gainsboro
RGB: 220, 220, 220
Hex: #DCDCDC
CMYK:
1 2 3 4
Corn
RGB: 251, 236, 93
Hex: #FBEC5D
CMYK:
1 2 3 4
Timberwolf
RGB: 219, 215, 210
Hex: #DBD7D2
CMYK:
Imperial
RGB: 104, 40, 96
Hex: #682860
CMYK:
1 2 3 4
1 2 3 4
Tardis Blue
RGB: 0, 59, 111
Hex: #003B6F
CMYK:
Rubine Red
RGB: 206, 0, 88
Hex: #CE0058
CMYK: 0, 100, 43, 12
1 2 3 4
1 2 3 4
Corn
RGB: 251, 236, 93
Hex: #FBEC5D
CMYK:
1 2 3 4
Timberwolf
RGB: 219, 215, 210
Hex: #DBD7D2
CMYK:
Notes: Remove American orange add hyperlink grey
Make corn the last for b135
All work in 134 imperial move later
Tardis doesn’t work in b136

## Slide 62
https://color.adobe.com/create/color-wheel
Analogous
F7AB19
D67F15
EE7624
D64615
F73119
Monochromatic
6E3610
F0A16C
EE7624
6E4931
BASB1C
Triad
A14508
4BED3B
EE7624
250CED
2010A1
Complementary
FF8C40
EE7624
0098A1
24E2ED
A1470C
Split Complementary
28A164
2FED8D
EE7624
0848A1
1871ED
Double Split Complementary
3BED93
EE7624
0C6AED
ED2F18
EDAC2F
Square
C7ED3B
EE7624
0CE1ED
A418ED
ED660C
Compound
87724A
EE7624
60EFCF
09BA61
BA7E09
Shades
6E3610
EE7624
FA7625
D46820
AD551A
Notes: These are the variants that work well with Fang Orange. The bolded items show well on the projectors.

## Slide 63
https://color.adobe.com/create/color-wheel
Analogous
85412D
8F3831
782f40
8F3176
792D85
Monochromatic
C44D69
8E626D
782F40
C48796
451B25
Triad
C43959
78743B
782F40
236178
43A2C4
Complementary
C46078
782F40
25C43D
2F783A
C43959
Split Complementary
93C460
577835
782F40
39C49D
297861
Double Split Complementary
5A783B
782F40
237860
6E2978
784435
Square
78683b
782F40
237830
293978
782337
Compound
DEBAAF
782F40
5D8555
70AB32
AB4D32
Shades
C44D69
782F40
853447
5E2532
38161E

## Slide 64
B135 Stage Right
A105
A105
B135 Stage Left
B135
B134, A105
B136